B20S
Handheld Thermography Camera

HIKMICRO B20S is designed for temperature measurement to find invisible faults. It’s equipped with a 256 x 192 resolution thermal detector and a 2MP optical detector. The thermography ranges from -20°C to 550°C (-4°F to 1022°F). Which meets most situations. It helps the staff to quickly find the fault according to accurate temperature measurements of high-temperature targets in the environment. Meanwhile, it provides assistance decisions and ensures safety. The device is mainly applied to various industries, such as building, HVAC, and auto industry, etc.

Key Features

- IR resolution: 256 x 192 (49 152 pixels), 2MP Visual Resolution

- Support SuperIR image enhancement technology

- NETD : ＜ 40 mK (@ 25 °C,F#=1.0)

- Temperature measurement range : -20°C to 550°C (-4°F to 1022°F）

- Accuracy: Max (±2°C/3.6°F, ±2%), for ambient temperature 15°C to 35°C (59°F to 95°F) and object temperature above 0°C (32°F)

- Measurement presets: Center spot, Hot spot, Cold spot, User Preset Point

- 25 Hz image frequency

- 3.2" 480 × 640 Resolution LCD Screen

- Up to 6 hours continuous running

| - Specification |
| --- |
| Infrared Image |  |
| IR Resolution | 256 × 192 (49,152 pixels) |
| SuperIR | Yes, on captured images and live view |
| NETD | < 40 mK (@ 25 °C, F#=1.0) |
| Image Frequency | 25 Hz |
| Detector Pitch | 12 μm |
| Spectral Range | 7.5 to 14 μm |
| Focal Length | 3.6 mm |
| F-number | F1.0 |
| Field of View (FOV) | 37.2° × 50.0° |
| Spatial Resolution (IFOV) | 3.3 mrad |
| Min. Focus Distance | 0.3 m (0.98 ft) |
| Focus Mode | Focus Free |
| Image Display |  |
| Visual Camera | 1600 × 1200 (2 MP) |
| Display | 480 × 640 Resolution, 3.2" LCD Screen |
| Color Alarm | Above alarm |
| Color Palettes | White Hot, Black Hot, Rainbow, Ironbow, Red Hot, Fusion |
| Image Modes | Thermal/Visual/Fusion/PIP |
| Measurement and Analysis |  |
| Object Temperature Range | -20°C to 550°C (-4°F to 1022°F) |
| Accuracy | Max. (± 2°C/3.6°F, ± 2%), for ambient temperature 15°C to 35°C (59°F to 95°F) and object temperature above 0°C (32°F) |
| Measurement Tools | Center Spot, Hot Spot, Cold Spot User-definable: 3 spots |
| Level and Span Mode | Auto/Manual |
| Data Storage and Communication |  |
| Storage Media | Built-in 16G eMMC (user available storage space is less than the nominal capacity) |
| Image Storage Capacity | Approx. 35,000 Images |
| Image File Format | Radiometric JPEG |
| General |  |
| Wi-Fi | 802.11 b/g/n (2.4GHz) |
| USB Interface | USB Type-C |
| LED Light | Yes |
| Power Supply | USB-C (5V/2A) |
| Battery Type | Rechargeable Li-ion battery |
| Battery Operating Time | Approx. 6 hours |
| Battery Charging Time | Approx. 3 hours fully charged |
| Protection Level | IP54 |
| Drop Test Height | 2 m (6.56 ft) |
| Working Temperature Range | -10°C to 50°C (14°F to 122°F) |
| Storage Temperature Range | -20°C to 60°C (-4°F to 140°F) |
| Relative Humidity | < 95% non-condensing |
| Weight | Approx. 380 g (0.84 lb) |
| Dimension | 221.7 × 73.5 × 123.8mm (8.73 × 2.89 × 4.87 in) |
| Tripod Mounting | UNC ¼”-20 |

- Available Model B20S

- Dimension

| - Optional Accessories |  |
| --- | --- |
| HM-B201-MACRO Thermography Camera Macro Lens | HM-2925ZJ-TM01-BRACKET Tripod Mounting Bracket | HM-B01-POUCH POUCH |

COMPLIANCE NOTICE: The thermal series products might be subject to export controls in various countries or regions, including without limitation, the United States, European Union, United Kingdom and/or other member countries of the Wassenaar Arrangement. Please consult your professional legal or compliance expert or local government authorities for any necessary export license requirements if you intend to transfer, export, re-export the thermal series products between different countries.